## Sheet1
CVE ID | Description | Weaknesses | baseScore | baseSeverity | attackVector | attackComplexity | privilegesRequired | userInteraction | confidentialityImpact | integrityImpact | availabilityImpact | exploitabilityScore | CVSS Version
CVE-2017-12723 | A Password in Configuration File issue was discovered in Smiths Medical Medfusion 4000 Wireless Syringe Infusion Pump, Version 1.1, 1.5, and 1.6. The pump stores some passwords in the configuration file, which are accessible if the pump is configured to allow external communications. | CWE-200 | 3.7 | LOW | NETWORK | HIGH | NONE | NONE | LOW | NONE | NONE | 2.2 | 3.0
CVE-2020-12045 | The Baxter Spectrum WBM (v17, v20D29, v20D30, v20D31, and v22D24) when used in conjunction with a Baxter Spectrum v8.x (model 35700BAX2), operates a Telnet service on Port 1023 with hard-coded credentials. | CWE-259; CWE-798 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.1
CVE-2014-5434 | Baxter SIGMA Spectrum Infusion System version 6.05 (model 35700BAX) with wireless battery module (WBM) version 16 has a default account with hard-coded credentials used with the FTP protocol. Baxter asserts no files can be transferred to or from the WBM using this account. Baxter has released a new version of the SIGMA Spectrum Infusion System, Version 8, which incorporates hardware and software changes. | CWE-259; CWE-798 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.0
CVE-2017-12724 | A Use of Hard-coded Credentials issue was discovered in Smiths Medical Medfusion 4000 Wireless Syringe Infusion Pump, Version 1.1, 1.5, and 1.6. The FTP server on the pump contains hardcoded credentials, which are not fully initialized. The FTP server is only accessible if the pump is configured to allow FTP connections. | CWE-798 | 8.1 | HIGH | NETWORK | HIGH | NONE | NONE | HIGH | HIGH | HIGH | 2.2 | 3.0
CVE-2023-0888 | An improper neutralization of directives in dynamically evaluated code vulnerability in the WiFi Battery embedded web server in versions L90/U70 and L92/U92 can be used to gain administrative access to the WiFi communication module. An authenticated user, having access to both the medical device WiFi network (such as a biomedical engineering staff member) and the specific B.Braun Battery Pack SP with WiFi web server credentials, could get administrative (root) access on the infusion pump communication module. This could be used as a vector to start further attacks | CWE-94; CWE-95 | 4.9 | MEDIUM | NETWORK | LOW | HIGH | NONE | NONE | HIGH | NONE | 1.2 | 3.1
CVE-2015-3954 | Hospira Plum A+ Infusion System version 13.4 and prior, Plum A+3 Infusion System version 13.6 and prior, and Symbiq Infusion System, version 3.13 and prior give unauthenticated users root privileges on Port 23/TELNET by default. An unauthorized user could issue commands to the pump. Hospira recommends that customers close Port 23/TELNET on the affected devices. Hospira has also released the Plum 360 Infusion System which is not vulnerable to this issue. | CWE-285 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.0
CVE-2014-5433 | An unauthenticated remote attacker may be able to execute commands to view wireless account credentials that are stored in cleartext on Baxter SIGMA Spectrum Infusion System version 6.05 (model 35700BAX) with wireless battery module (WBM) version 16, which may allow an attacker to gain access the host network. Baxter has released a new version of the SIGMA Spectrum Infusion System, Version 8, which incorporates hardware and software changes. | CWE-255; CWE-312 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.0
CVE-2014-5432 | Baxter SIGMA Spectrum Infusion System version 6.05 (model 35700BAX) with wireless battery module (WBM) version 16 is remotely accessible via Port 22/SSH without authentication. A remote attacker may be able to make unauthorized configuration changes to the WBM, as well as issue commands to access account credentials and shared keys. Baxter asserts that this vulnerability only allows access to features and functionality on the WBM and that the SIGMA Spectrum infusion pump cannot be controlled from the WBM. Baxter has released a new version of the SIGMA Spectrum Infusion System, Version 8, which incorporates hardware and software changes. | CWE-287; CWE-592 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.0
CVE-2017-12722 | An Out-of-bounds Read issue was discovered in Smiths Medical Medfusion 4000 Wireless Syringe Infusion Pump, Version 1.1, 1.5, and 1.6. A third-party component used in the pump reads memory out of bounds, causing the communications module to crash. Smiths Medical assesses that the crash of the communications module would not impact the operation of the therapeutic module. | CWE-125 | 5.3 | MEDIUM | NETWORK | LOW | NONE | NONE | NONE | NONE | LOW | 3.9 | 3.0
CVE-2022-43557 | The BD BodyGuard™ infusion pumps specified allow for access through the RS-232 (serial) port interface. If exploited, threat actors with physical access, specialized equipment and knowledge may be able to configure or disable the pump. No electronic protected health information (ePHI), protected health information (PHI) or personally identifiable information (PII) is stored in the pump. | CWE-1299; CWE-287 | 5.3 | MEDIUM | PHYSICAL | HIGH | NONE | NONE | LOW | LOW | HIGH | 0.5 | 3.1
CVE-2017-12725 | A Use of Hard-coded Credentials issue was discovered in Smiths Medical Medfusion 4000 Wireless Syringe Infusion Pump, Version 1.1, 1.5, and 1.6. The pump with default network configuration uses hard-coded credentials to automatically establish a wireless network connection. The pump will establish a wireless network connection even if the pump is Ethernet connected and active; however, if the wireless association is established and the Ethernet cable is attached, the pump does not attach the network stack to the wireless network. In this scenario, all network traffic is instead directed over the wired Ethernet connection. | CWE-798 | 5.6 | MEDIUM | NETWORK | HIGH | NONE | NONE | LOW | LOW | LOW | 2.2 | 3.0
CVE-2020-12041 | The Baxter Spectrum WBM (v17, v20D29, v20D30, v20D31, and v22D24) telnet Command-Line Interface, grants access to sensitive data stored on the WBM that permits temporary configuration changes to network settings of the WBM, and allows the WBM to be rebooted. Temporary configuration changes to network settings are removed upon reboot. | CWE-732 | 9.4 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | LOW | HIGH | 3.9 | 3.1
CVE-2017-12720 | An Improper Access Control issue was discovered in Smiths Medical Medfusion 4000 Wireless Syringe Infusion Pump, Version 1.1, 1.5, and 1.6. The FTP server on the pump does not require authentication if the pump is configured to allow FTP connections. | CWE-306 | 8.1 | HIGH | NETWORK | HIGH | NONE | NONE | HIGH | HIGH | HIGH | 2.2 | 3.0
CVE-2014-5403 | Hospira MedNet before 6.1 uses hardcoded cryptographic keys for protection of data transmission from infusion pumps, which allows remote attackers to obtain sensitive information by sniffing the network. | CWE-310; CWE-321 | 6.8 | MEDIUM | NETWORK | LOW | SINGLE | NONE | COMPLETE | NONE | NONE | 8 | 2.0
CVE-2017-12726 | A Use of Hard-coded Password issue was discovered in Smiths Medical Medfusion 4000 Wireless Syringe Infusion Pump, Version 1.1, 1.5, and 1.6. Telnet on the pump uses hardcoded credentials, which can be used if the pump is configured to allow external communications. Smiths Medical assesses that it is not possible to upload files via Telnet and the impact of this vulnerability is limited to the communications module. | CWE-798 | 7.3 | HIGH | NETWORK | LOW | NONE | NONE | LOW | LOW | LOW | 3.9 | 3.0
CVE-2014-5431 | Baxter SIGMA Spectrum Infusion System version 6.05 (model 35700BAX) with wireless battery module (WBM) version 16 contains a hard-coded password, which provides access to basic biomedical information, limited device settings, and network configuration of the WBM, if connected. The hard-coded password may allow an attacker with physical access to the device to access management functions to make unauthorized configuration changes to biomedical settings such as turn on and off wireless connections and the phase-complete audible alarm that indicates the end of an infusion phase. Baxter has released a new version of the SIGMA Spectrum Infusion System, version 8, which incorporates hardware and software changes. | CWE-259; CWE-798 | 6.8 | MEDIUM | PHYSICAL | LOW | NONE | NONE | HIGH | HIGH | HIGH | 0.9 | 3.0
CVE-2020-12043 | The Baxter Spectrum WBM (v17, v20D29, v20D30, v20D31, and v22D24) when configured for wireless networking the FTP service operating on the WBM remains operational until the WBM is rebooted. | CWE-672 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.1
CVE-2015-3459 | The communication module on the Hospira LifeCare PCA Infusion System before 7.0 does not require authentication for root TELNET sessions, which allows remote attackers to modify the pump configuration via unspecified commands. | CWE-264 | 10 | HIGH | NETWORK | LOW | NONE | NONE | COMPLETE | COMPLETE | COMPLETE | 10 | 2.0
CVE-2017-12721 | An Improper Certificate Validation issue was discovered in Smiths Medical Medfusion 4000 Wireless Syringe Infusion Pump, Version 1.1, 1.5, and 1.6. The pump does not validate host certificates, leaving the pump vulnerable to a man-in-the-middle (MITM) attack. | CWE-295 | 5.9 | MEDIUM | NETWORK | HIGH | NONE | NONE | HIGH | NONE | NONE | 2.2 | 3.0
CVE-2020-12039 | Baxter Sigma Spectrum Infusion Pumps Sigma Spectrum Infusion System v's6.x model 35700BAX & Baxter Spectrum Infusion System v's8.x model 35700BAX2 contain hardcoded passwords when physically entered on the keypad provide access to biomedical menus including device settings, view calibration values, network configuration of Sigma Spectrum WBM if installed. | CWE-259; CWE-798 | 2.4 | LOW | PHYSICAL | LOW | NONE | NONE | LOW | NONE | NONE | 0.9 | 3.1
CVE-2015-3956 | Hospira Plum A+ Infusion System version 13.4 and prior, Plum A+3 Infusion System version 13.6 and prior, and Symbiq Infusion System, version 3.13 and prior accept drug libraries, firmware updates, pump commands, and unauthorized configuration changes from unauthenticated devices on the host network. Hospira recommends that customers close Port 20/FTP and Port 23/TELNET on the affected devices. Hospira has also released the Plum 360 Infusion System which is not vulnerable to this issue. | CWE-345 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.0
CVE-2020-12047 | The Baxter Spectrum WBM (v17, v20D29, v20D30, v20D31, and v22D24), when used with a Baxter Spectrum v8.x (model 35700BAX2) in a factory-default wireless configuration enables an FTP service with hard-coded credentials. | CWE-259; CWE-798 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.1
CVE-2020-12040 | Sigma Spectrum Infusion System v's6.x (model 35700BAX) and Baxter Spectrum Infusion System Version(s) 8.x (model 35700BAX2) at the application layer uses an unauthenticated clear-text communication channel to send and receive system status and operational data. This could allow an attacker that has circumvented network security measures to view sensitive non-private data or to perform a man-in-the-middle attack. | CWE-319 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.1
CVE-2017-12718 | A Classic Buffer Overflow issue was discovered in Smiths Medical Medfusion 4000 Wireless Syringe Infusion Pump, Version 1.1, 1.5, and 1.6. A third-party component used in the pump does not verify input buffer size prior to copying, leading to a buffer overflow, allowing remote code execution on the target device. The pump receives the potentially malicious input infrequently and under certain conditions, increasing the difficulty of exploitation. | CWE-119; CWE-120 | 8.1 | HIGH | NETWORK | HIGH | NONE | NONE | HIGH | HIGH | HIGH | 2.2 | 3.0