## Sheet1
CVE ID | Description | Weaknesses | baseScore | baseSeverity | attackVector | attackComplexity | privilegesRequired | userInteraction | confidentialityImpact | integrityImpact | availabilityImpact | exploitabilityScore | CVSS Version | exposure_tags
CVE-2014-5431 | Baxter SIGMA Spectrum Infusion System version 6.05 (model 35700BAX) with wireless battery module (WBM) version 16 contains a hard-coded password, which provides access to basic biomedical information, limited device settings, and network configuration of the WBM, if connected. The hard-coded password may allow an attacker with physical access to the device to access management functions to make unauthorized configuration changes to biomedical settings such as turn on and off wireless connections and the phase-complete audible alarm that indicates the end of an infusion phase. Baxter has released a new version of the SIGMA Spectrum Infusion System, version 8, which incorporates hardware and software changes. | CWE-259; CWE-798 | 6.8 | MEDIUM | PHYSICAL | LOW | NONE | NONE | HIGH | HIGH | HIGH | 0.9 | 3.0 | Keypad/physical, Wi-Fi/wireless
CVE-2014-5432 | Baxter SIGMA Spectrum Infusion System version 6.05 (model 35700BAX) with wireless battery module (WBM) version 16 is remotely accessible via Port 22/SSH without authentication. A remote attacker may be able to make unauthorized configuration changes to the WBM, as well as issue commands to access account credentials and shared keys. Baxter asserts that this vulnerability only allows access to features and functionality on the WBM and that the SIGMA Spectrum infusion pump cannot be controlled from the WBM. Baxter has released a new version of the SIGMA Spectrum Infusion System, Version 8, which incorporates hardware and software changes. | CWE-287; CWE-592 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.0 | SSH, Wi-Fi/wireless
CVE-2014-5433 | An unauthenticated remote attacker may be able to execute commands to view wireless account credentials that are stored in cleartext on Baxter SIGMA Spectrum Infusion System version 6.05 (model 35700BAX) with wireless battery module (WBM) version 16, which may allow an attacker to gain access the host network. Baxter has released a new version of the SIGMA Spectrum Infusion System, Version 8, which incorporates hardware and software changes. | CWE-255; CWE-312 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.0 | Cleartext, Wi-Fi/wireless
CVE-2014-5434 | Baxter SIGMA Spectrum Infusion System version 6.05 (model 35700BAX) with wireless battery module (WBM) version 16 has a default account with hard-coded credentials used with the FTP protocol. Baxter asserts no files can be transferred to or from the WBM using this account. Baxter has released a new version of the SIGMA Spectrum Infusion System, Version 8, which incorporates hardware and software changes. | CWE-259; CWE-798 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.0 | FTP, Wi-Fi/wireless
CVE-2020-12039 | Baxter Sigma Spectrum Infusion Pumps Sigma Spectrum Infusion System v's6.x model 35700BAX & Baxter Spectrum Infusion System v's8.x model 35700BAX2 contain hardcoded passwords when physically entered on the keypad provide access to biomedical menus including device settings, view calibration values, network configuration of Sigma Spectrum WBM if installed. | CWE-259; CWE-798 | 2.4 | LOW | PHYSICAL | LOW | NONE | NONE | LOW | NONE | NONE | 0.9 | 3.1 | Keypad/physical
CVE-2020-12040 | Sigma Spectrum Infusion System v's6.x (model 35700BAX) and Baxter Spectrum Infusion System Version(s) 8.x (model 35700BAX2) at the application layer uses an unauthenticated clear-text communication channel to send and receive system status and operational data. This could allow an attacker that has circumvented network security measures to view sensitive non-private data or to perform a man-in-the-middle attack. | CWE-319 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.1 | Cleartext
CVE-2020-12041 | The Baxter Spectrum WBM (v17, v20D29, v20D30, v20D31, and v22D24) telnet Command-Line Interface, grants access to sensitive data stored on the WBM that permits temporary configuration changes to network settings of the WBM, and allows the WBM to be rebooted. Temporary configuration changes to network settings are removed upon reboot. | CWE-732 | 9.4 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | LOW | HIGH | 3.9 | 3.1 | Telnet
CVE-2020-12043 | The Baxter Spectrum WBM (v17, v20D29, v20D30, v20D31, and v22D24) when configured for wireless networking the FTP service operating on the WBM remains operational until the WBM is rebooted. | CWE-672 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.1 | FTP, Wi-Fi/wireless
CVE-2020-12045 | The Baxter Spectrum WBM (v17, v20D29, v20D30, v20D31, and v22D24) when used in conjunction with a Baxter Spectrum v8.x (model 35700BAX2), operates a Telnet service on Port 1023 with hard-coded credentials. | CWE-259; CWE-798 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.1 | Telnet
CVE-2020-12047 | The Baxter Spectrum WBM (v17, v20D29, v20D30, v20D31, and v22D24), when used with a Baxter Spectrum v8.x (model 35700BAX2) in a factory-default wireless configuration enables an FTP service with hard-coded credentials. | CWE-259; CWE-798 | 9.8 | CRITICAL | NETWORK | LOW | NONE | NONE | HIGH | HIGH | HIGH | 3.9 | 3.1 | FTP, Wi-Fi/wireless